| Bauer Core Standard Protocol |
| --- |
| Title: AutoGenprep 965 Tissue Protocol |
| Pages: 2 | Revision: 1.3 | Date: 09/26/2014 |
| Author(s): Claire Reardon | Reviewers: Christian Daly |
| Contact: claire@cgr.harvard.edu | Comment: |

1. Purpose

This protocol describes how to extract genomic DNA from tissue using the Mouse Tail DNA protocol on the AutoGenprep 965.  The protocol is designed as a reference and is not a substitute for training.  Users must complete a training session before using any of the Bauer Core’s instrumentation.

2. Materials

2.1.  up to 50mg of tissue per sample (up to 384 samples in the deep well plates below)

2.2.  2ml 96 Deep Well Plate(s) (Corning # 3960)

2.3.  Tissue DNA Extraction Kit (Autogen # AGKT965TD)

2.4.  Plate sealers (eg. Hypertask # MPCAL096)

3. Instrumentation - AutoGenPrep 965

4. Reagent preparation

4.1  Prepare Protinase K solution (ProK/M1/M2)

4.1.1.  Dissolve the protinase K in solution M1.

4.1.1.1.  The concentration will depend on how hard the tissue is to digest.

Use 1mg/ml for tissues that are difficult to digest.

Use 0.4 mg/ml for easier tissues.

4.1.2.  Mix the protinase K/M1 solution 1:1 with M2 solution.

5. Procedure

5.1  Digest the tissue

5.1.1.  This step should be optimized for each tissue type to achieve full digestion.

5.1.2.  Place up to 50mg of tissue into each well of a 96 deep well plate

5.1.3.  Add from 300ul to 500ul of the ProK/M1/M2 solution to each well

5.1.3.1. The plates will be centrifuged and must be balanced.  If the plates

are not full, split the samples into the upper rows of two plates.

5.1.3.  Digest at 55-60ºC from 1h to 16h (until the solution is clear)

5.2.  Load the Instrument

5.2.1.  Make sure the waste containers are empty.

5.2.2.  Place the deep well plates containing the samples onto the sample rack.

The sample rack is on the right hand side of the deck

Well A1 should be in the back left corner.

5.2.3.  Place an empty deep well plate to the left of each sample plate to hold the

purified DNA.

5.2.4.  Place tip boxes to the right of the sample plates.

Load one box of tips for reagent dispensing (rear tip box location)

Load one box of tips for each sample plate.

5.2.5.  Load the Reagent Reservoirs according to the table below:

|  |  | Vol/Sample | Vol. | Per | Sample # |
| --- | --- | --- | --- | --- | --- |
| Reagent | Reservoir | (ml) | 96 | 192 | 288 | 384 |
| R3 Deproteinization | 3 | 0.18 | 42 | 60 | 77 | 94 |
| R4 DNA Precipitation | 4 | 0.20 | 45 | 64 | 83 | 102 |
| R5/6/7 DNA Wash | 5,6,7 | 0.45 | 69 | 112 | 155 | 198 |
| R8 Digestion | 8 | 0.05 | 30 | 35 | 40 | 45 |
| R9 DNA Resuspension | 9 | 0.05 | 30 | 35 | 40 | 45 |

5.2.5.1.  See table for reagent volumes for 96, 192, 288 or 384 samples.

5.2.5.2.  For other numbers of samples, multiply Vol/Sample by the number of

samples and add 25 ml of “dead” volume.

5.2.5.3.  Reagent 5/6/7 can be used for up to 3 DNA washes

Fill one reservoir for each wash desired (default = 2).

5.3.  Run the Software

5.3.1.  Turn on the Autogen.

5.3.2.  Open the Autogen software.

5.2.2.1.  Select On Line and hit Ok.

5.3.3.  Choose to run a Standard Protocol.

5.3.3.1.  Check all the items in the checklist to make sure they are in place.

5.3.3.2.  Select the protocol to run – Mouse Tail.

5.3.3.3.  Select Automated Run.

5.3.4.  Choose either the Standard protocol or create a User protocol.

5.3.4.1.  Standard protocols can’t be edited except for sample number.

5.3.4.2.  Custom protocols can be edited at any step.

Commonly, users change the # of washes or resuspension vol.

5.3.4.3.  Click OK when the protocol is edited as desired.

5.3.5.  Click Start to run the System check.

5.3.6.  Click OK to begin.

5.4.  Clean Up

5.4.1.  Dispose of all tips and used plates in biohazard waste.

5.4.2.  Empty waste bin 1 and dispose of the contents as a biohazard.

5.4.3.  Empty leftover reagents in the reagent trays into waste bin 2.

5.4.3.1.  Rinse the reagent trays with water.

5.4.4. Empty waste bin 2 and dispose of the contents as chemical waste.

5.4.4.1.  Label the waste as toxic and flammable.

5.4.4.2.  Record the following waste composition:

Butyl Alcohol 	5%

Chloroform	5%

Isopropanol	20%

Phenol		5%

Ethanol	34%

5.4.5.  Shut down the software

5.4.6.  Turn off the power to the machine.

5.4.7.  Log off the computer.